HOW  TO SET UP A TC307/TCR110 WITH A LASER PLUMB OVER A POINT

This tutorial will show you how to set TC307/TCR110 with a laser plumb over a point.

It takes approximately  10 mins

# You will need:

- TC307/TCR110 with laser plumb

- Tripod to suit the instrument

- Fixed point on the ground = nail

- Set tripod at an appropriate height approximately over the nail.

- Put the instrument on the tripod and clamp the theodolite with the instrument screw.

- Turn on the instrument.

- Turn on the laser pointer

- Turn the footscrews until the laser centres on the nail.

#### With your foot on the foot of the tripod leg to prevent movement, adjust each tripod leg in

turn until the circular bubble on the instrument is level in all directions.

Note: It is most important to ensure that no movement of the tripod feet takes place

Checkpoint 1:	Check that the laser is still within 5mm of the nail.  If not return to 4.

- Turn the instrument until the front electronic bubble displayed is parallel to the line of two

footscrews.

- Turn the footscrews under the front bubble until the bubble is level.

- Level the other electronic bubble using the 3rd footscrew only.  Do not move the previous

2 footscrews.

- Slacken the instrument screw 1 turn.

- While watching the laser spot, slide the instrument until the laser spot is over the nail.

Note:	Try not to rotate the instrument while you do this movement as it will put the

instrument off level.

- Do 8, 9, 10, 11 again.

Checkpoint 2:	  Turn the instrument 180º and check that laser is still centred on nail.  If the laser spot does not remain centred on the nail or if the electronic bubbles move off level the instrument is out of adjustment, take it to your service workshop.

GLOSSARY:

Instrument Screw:	The screw on top of the tripod for holding on the instrument

Circular Bubble:	The circular bubble on the instrument for rough levelling adjustment.

Electronic Bubble:	The levelling bubble shown on the instrument display

Tripod leg:		That part of the tripod from the head to the foot.

Tripod foot:		The lowest part of the tripod that you put your own foot on to push it into the ground, also has a point on the end.

TROUBLESHOOTING:

The laser spot seems blocked:

There is a hanger for plumb bobs inside some instrument screws, push it to one side and there should be a holder to keep it out of the way.

The laser spot is too bright/too dim:

There is an intensity adjustment for the laser spot, see it on screen and adjust accordingly.

There is not sufficient adjustment on the footscrews to bring the laser spot over the nail:

Reposition the tripod and start at 1 again.

The electronic bubble will not stay centred when I do Checkpoint 2, even though I have repeated from 6 several times:

The instrument is out of adjustment.  Take the instrument to your service workshop for adjustment.

The laser spot does not stay centred when I do Checkpoint 2:

The instrument is out of adjustment. Take the instrument to your service workshop for adjustment.